Cell suspension as a model system for electrochemical analysis

Solveg Kovač, Romina Kraus, Sunčana Geček  and Vera Žutić*

Ruđer Bošković Institute, Center for Marine and Environmental Research, P.O.B. 1016, 10001 Zagreb, Croatia

Cell suspension model for electroanalysis

Key words: attachment signals, dropping mercury electrode, Dunaliella tertiolecta, organic particle analysis

ABSTRACT

Cell suspensions of marine phytoplankton Dunaliella tertiolecta in 0.1 M NaCl solution are proposed as a suitable model monodispersed system (particle diameter 6-10(m) for the calibration of electrochemical response in natural aquatic samples containing organic surface-active particles. The electrochemical analysis is performed by direct recording of chronoamperometric curves of oxygen reduction in suspensions using a fast dropping mercury electrode. The electrode acts as the adhesion sensor. The adhesion of individual surface-active particles suspended in natural seawater, analogously to Dunaliella tertiolecta cells in model suspensions, results in well resolved attachment signals in amperometric curves. The calibration curve presents dependence of attachment frequency on cell densities in the concentration range 106 to 2.5x107 particles L-1 and it can be used for the determination of particle abundance in a seawater sample. The advantage of the electrochemical approach over more conventional methodologies for particle analysis is discussed.

INTRODUCTION

A standard way of characterization of  marine organic matter in seawater and freshwater samples is either measurement of “dissolved” or “particulate” organic carbon (after filtration, 0.45 (m pore size),1,2 or determination of individual compounds or their classes, after a complex pretreatment of a water sample.3

Difficulties in physico-chemical speciation of organic constituents of seawater stems in part from low concentrations involved (total organic carbon is 0.3-3 mg C L-1) and from complex nature of organic matter. The classes of compounds present range from the simplest of hydrocarbons to the most complex biogeopolymers. Concentrations of carbohydrates, organic acids, proteins, lipids and other identified substances explain not more than 10% of the organic carbon in seawater.4

The electrochemical characterization of surfactant activity of organic constituents in seawater using a simple  polarographic technique, first introduced by Zvonarić et al. (1973),5,6 enabled identification of fluid surface-active particles,7-9 a highly reactive class of organic particles that was not amenable to analysis by conventional methods. Surface-active particles could be detected electrochemically in the size range (1(m. Their concentration in coastal and estuarine waters was estimated to be in the concentration range 106 – 109 particles L-1.

In recent years there has been a dramatic change in our knowledge of particulate matter in the ocean because of the discovery of new classes of highly abundant organic particles: colloids,10,11 submicron particles,12 and transparent microscopic particles13 which had remained undetected by previous techniques. These range in size from submicron to hundreds of microns  and  depending on the size range their  reported concentrations vary from 105 to 1014 particles L-1. These particles play major role in the ocean’s ecology and chemistry.

A major characteristics of a water sample containing colloidal particles is its intrinsic instability due to continuing aggregation processes and microbial activity.  Consequently, sampling and sample processing should be shortned and simplified as much as possible.14,15 The electrochemical particle analysis, being direct rapid and simple, meets these requirements and also offers the possibility of single particle analysis. One of still unresolved problems is a proper calibration of the electrochemical response. Here we investigate a simple biological standard for calibration.

THE ELECTROCHEMICAL METHOD

The electrochemical technique employed is a modification of a widely used polarographic technique for measuring surface-active constituents in environmental aquatic samples.16 The method is based on chronoamperometric measurement of modification of the interfacial turbulence during oxygen reduction at a fast dropping mercury electrode in aqueous electrolyte solutions.17-20 Attachment and adhesion of organic particles causes a transient increase in the interfacial turbulence resulting in spike shaped attachment signals of individual particles.21-25 The dropping mercury electrode has a fast growing and renewable surface, and an experiment can be repeated many times in the same environment.

##### EXPERIMENTAL

# Cell culture and cell suspensions

We used laboratory cultures of marine nanoflagellate Dunaliella tertiolecta Butcher (strain CCMP 1320), obtained from Provasoli-Guillard Center for Culture of Marine Phytoplankton, Bigelow Laboratory for Ocean Sciences as a source of model particles. Dunaliella tertiolecta cells do not possess cell wall, only a flexible outer membrane (Figure 1.). The size of cells is in the range 6-10 (m.

The cells were grown in F-2 medium (composed of nutrient salts, essential trace metals and vitamins)26 in batch cultures at ambient conditions. The medium was prepared by adding F-2 nutrients with Sterile Acrodisc (0.2 (m, Gelman Sciences) to sterilized seawater. Cell densities in culture, in stock and in the analyzed suspensions were counted using a light microscope (Hund Wetzlar H 500) with haemacytometer (Fuchs-Rosenthal, Fein-Optik Jena, Germany, Tiefe 0.2 mm).

After 6-8 days of growth, cell density in the medium reached up to 109 L-1. Then, the cells were separated from the growth medium using mild centrifugation (1500g, 5 minutes), the supernatant was carefully removed, and the loose pelet washed with filtered seawater. This procedure was repeated 2 times to remove traces of the growth medium. The pelet was then resuspended in 2 ml of filtered seawater which served as a stock suspension. Cell densities in stock suspensions were 1-6x1010 L-1.

Aliquotes of stock suspension were added with a micropipette to 25 ml of the organic-free electrolyte (0.1M NaCl) to prepare suspensions of given cell densities immediately before electrochemical measurement. pH in measured suspensions was 8.2 and mantained constant with  addition of 5x10-3M NaHCO3. Water used for the preparation of organic-free electrolyte or for dilution of seawater samples was ultra-pure MilliQ water. The purity of the system was controlled by measuring the polarographic maximum of oxygen reduction.

## Electrochemical measurement

Figure 2. shows a block scheme of the electrochemical measuring system. Chronoamperometric measurements were performed using a PAR 174A Polarographic Analyzer. The current-time curves (time dependence of the instantaneous current during one drop life)27 at the constant potential -400 mV (versus a Ag/AgCl reference electrode as described below) were recorded and stored using a Nicolet 3091 digital oscilloscope connected to a PC.

Measurements were performed in a standard Methrom vessel with 25 ml of a freshly prepared cell suspension, termostated at 20(1(C. The measured samples were saturated with air and the vessel was open to air throughout the experiments.

A fast dropping mercury electrode with drop life 2.0 s, flow rate 6.03 mg s-1 and maximum surface area 4.57 mm2 (Figure 3) was used with Ag/AgCl electrode as a reference in the three-electrode configuration. The reference electrode was separated from the measured suspension by a ceramic frit and its potential in 0.1 M NaCl was  + 2mV vs. calomel electrode (1 M KCl).

### RESULTS AND DISCUSSION

All experiments were performed using  0.1 M NaCl as supporting electrolyte  because  the electrochemical literature28 contains reliable physical and chemical information about the mercury electrode/0.1 M NaCl interface. Besides, in 0.1 M NaCl solutions the streaming maximum of oxygen reduction is well pronounced and has been studied before.20 At a potential of –400 mV, the charge density of the hydrophobic mercury surface is +3.8 (C cm-2, and the interfacial tension is close to electrocapillary maximum.29 Under such conditions at the interface, hydrophobic and electrostatic attractions are expected to prevail in the adhesion phenomena of natural organic particles.

Marine microorganism Dunaliella tertiolecta was chosen as model particle because of its suitable size and membrane properties. It is easily available, simple to grow in batch culture, and it forms stable suspensions of single cells that approach characteristics of a monodispersed system. Concentrations of cells in suspensions can be dosed and measured precisely. Suspensions are easy to prepare by adding aliquots of stock cell suspension to the electrolyte solution immediately prior to the measurements.

Figure 4a shows current-time curves of oxygen reduction in a seawater sample from Northen Adriatic30. Characteristic electrical signals appearing as sharp spikes on I-t curves are the result of random attachment of surface-active particles onto the electrodes. The attachment signals appear at irregular intervals and with different amplitudes.

To analyze the recorded current-time curves in terms of particle abundance in the sample we conducted the series of calibrating experiments using suspensions of Dunaliella tertiolecta cells under identical experimental conditions.

A typical response of a Dunaliella tertiolecta cells suspension is given in Figure 4b. The cell concentration was 6.8x106 L-1. The attachment signals of similar shape and amplitude appearing on I-t curves are the result of random attachment of individual cells. In Figure 5 we present examples of I-t curves recorded over one drop life with a higher time resolution for a better observation of the attachment signals of individual cells. The attachment signals are well defined, they appear single or in a sequence separated from each other, and are very similar in shape. Their amplitudes vary between 1.2 and 1.8 (A and duration between 100 and 140 ms. It has been established in previous studies performed in the laboratory that in each attachment signal the sharp increase of current reflects initial molecular contact between a cell and the mercury surface, while its slow decay follows spreading of the cell material over the surface after rupture of the cell membrane.24,31,32

For the series of increasing cell densities in the range from 2.3x106 to 3.1x108 L-1, the frequency of appearance of attachment signals is expressed as number of attachment signals per drop life and plotted as a function of cell density (Figure 6). The attachment frequencies were presented as mean values with standard deviation, obtained by analyzing of 30 I-t curves. The values of standard deviations reflect the stochastic nature of the process.32 In the concentration range where the mean frequency of attachment signals increased proportionally, the individual attachment signals are well resolved and do not overlap with each other. This concentration range, 2.3x106 to 108 L-1, is best suited for the construction of a calibration curve. With further increase of cell concentration in the suspension a situation arises where at a given time instant there is more than one cell attaching to the electrode surface resulting in a sporadic overlap of attachment signals or a significant reduction of the electrode free surface area32. Counting of individual attachment signals becomes difficult when the attachment frequency approaches 20 and this represents the limit of measurable range (Figure 7).

In order to obtain information about the probabilities of appearance of a given number of signals during a drop-life, detailed analysis of attachment signals were performed in the sequence of I-t curves. Figure 8 shows frequency distribution for three representative cell concentrations, 2.4x106, 9.5x106 and 2.4x107 L-1. The range of the numbers of signals appearing over a drop-life increases with increasing cell density. A hypothesis test (χ2 – test) has been introduced to decide how reasonable it is to assume that a Poisson probability model fits particular data sets. With respect to the fact that χ2- approximation is adequate if no expected frequency is less than five33 and by taking into account the fact that data sets were used to estimate the Poisson’s means, the χ2-distributions against which the statistics are to be tested have 1, 3 and 5 degrees of freedom respectively for the three cell densities considered. The values obtained for the χ2-test statistic are 0.034, 6.728 and 3.998, meaning that significance probabilities are 0.85, 0.08 and 0.54 respectively. Therefore, we can conclude that there is little evidence to reject the null hypothesis that data may be fitted by Poisson’s distribution33 which is typical of rear events.

To obtain the best results, knowing the nature of the process, the experiments were repeated by counting attachment signals on 100 I-t curves for each cell density. The results are plotted in Figure 9. This calibration curve can then be used for the determination of concentrations of surface-active particles in natural seawater samples, as illustrated below.

The number of attachment signals obtained by analyzing I-t curves is exemplified in Figure 4a was 43 for 100 I-t curves. According to the calibration curve the corresponding concentration of surface-active particles is 1.5x106 L-1. As the measured sample was diluted 1:5, the concentration in the original undiluted seawater is then 7.5x106 L-1. This concentration refers to surface-active particles in the size range ( 1 (m. With the simple electrochemical equipment used the attachment signals for smaller particles (( 1 (m) cannot be distinguished from the instrumental noise.21, 24,34,35

The optimum measurement range under the experimental conditions used in this work is 2x106 to 108 particles L-1.  This range can be extended to lower concentrations by prolonged recording of I-t curves (duration of analysis ( 5 minutes) and also to higher particle concentration by appropriate dilution with 0.1 M NaCl solution. For field analysis when a large number of samples is involved a procedure for direct measurements in undiluted seawater can be easily adopted.

#### GENERAL CONCLUSIONS

The unique advantage of the electrochemical approach over impedance volume measurements, fluorescence flow cytometry and microscopic observation of single particles generally used in characterization of aquatic particles14 is:

- the possibility to selectively analyze aqueous samples directly without fractionally pretreatment

- the  inexpensive equipment and simple manipulation

- simple model system for calibration.

The method is suitable for selective analysis of surface-active organic particles in the size range 1-100 (m16 and the concentration range 105-107 particles L-1.

In future investigations Dunaliella tertiolecta model system can be used to characterize the size of particles (according the amplitude of attachment signals)34 and also the sticking characteristics of natural particles. Moreover, it would be possible to develop more sophisticated methods of signal treatment and use the electrochemical method for a number of industrial applications, such as to control the purity of water for specific purposes.36,37

ACKNOWLEDGEMENTS

This research was sponsored by the Ministry of Science and Technology of Croatia, Project P-1508.

FIGURE CAPTIONS

Figure 1. Electron micrograph of a thin section through a Dunaliella tertiolecta cell, showing the cell membrane and details of cell structure (magnification 18.000:1). Courtesy of dr. Mercedes Wrischer.

Figure 2. Block scheme of the electrochemical measuring system

Figure 3. Periodic change of the surface area of the dropping mercury electrode used in this work

Figure 4. Current-time curves of oxygen reduction recorded in an Adriatic seawater sample, North Adriatic, station 103,30 depth 10m, February 27 1998, (A); and in cell suspension of 6.8x106 cells L-1 Dunaliella tertiolecta in 0.1 M NaCl (B).

Actual recordings; potential –400 mV, time resolution 10 ms per point. Prior to measurement the seawater sample was diluted (1:5) with MilliQ water to reach the ionic strength of a 0.1 M NaCl solution.

Figure 5. Examples of current-time curves recorded with time resolution of 2 ms per point in the suspension of 4.6x106 cells L-1  in 0.1 M NaCl.

Figure 6. Dependence of mean attachment frequency on cell densities in the suspensions of Dunaliella tertiolecta cells in the concentration range 2.3x106-3.1x108 cells L-1.

Figure 7. Typical current-time curves recorded in the suspension of 3.1x108 cells L-1 in 0.1 M NaCl; time resolution 2 ms per point

Figure 8. Frequency dependence of the number of attachment signals per I-t curve recorded at –400 mV in suspensions of Dunaliella tertiolecta cells for 60 I-t curves analyzed in 0.1 M NaCl solutions. Cell densities were 2.4x106, 9.5x106 and 2.4x107 cells L-1.

Figure 9. The calibration curve: frequency of attachment signals on 100 consecutive

I-t vs cell density in suspensions of Dunaliella tertiolecta cells.

REFERENCES

1. E.D.Goldberg, M.Baker and D.L.Fox, J.Mar.Res. 11 (1952) 197-202.

2. E.R.Duursma and R.Dawson (eds.), Marine Organic Chemistry, Elsevier, Amsterdam,  1982, p 456.

3. P.J.Wangersky, Mar.Chem. 41 (1993) 61-74.

4. P.J.Wangersky in T.R.P. Gibb (ed.), Analytical Methods in Oceanography, Adv.                            Chem. Ser., Washington 147 (1975) 148-162.

- T.Zvonarić, V.Žutić and M.Branica, Thalas. Jugosl. 9 (1973) 65-73.

- T.Zvonarić, M. Sc. Thesis, Electrochemical Determination of Surface Active Substances in Seawater, University of Zagreb, 1975. (in Croatian)

- V.Žutić, T.Pleše, J.Tomaić and T.Legović, Mol. Cryst. Liq. Cryst. 113 (1984) 131-145.

- V.Žutić and J.Tomaić, Mar. Chem. 23 (1988) 5167.

- V.Žutić and T.Legović, Nature 328 (1987) 612-614.

- M.L.Wells and E.D.Goldberg, Nature 353 (1991) 342-344.

- M.L.Wells, Nature 391 (1998) 530-531.

- I.Koike, H.Shigemitsu, T.Kazuki and K.Kogure, Nature 345 (1990) 242-244.

- A.L.Alldredge, U.Passow and B.E.Logan, Deep Sea Res. 40 (1993) 1131-1140.

- J.Buffle, D.Perret, M.Newman in J.Buffle and H.P. van Leeuwen (eds.), Environmental Particles, Environmental Analytical and Physical Chemistry Series, Vol.1, Lewis Publishers, Boca Raton, 1992, p. 171-220.

- J.Buffle and G.G.Leppard, Environ. Sci. Technol. 29 (1995) 2176-2184.

- V.Žutić, V.Svetličić and J.Tomaić, Pure Appl.Chem. 62 (1990) 2269-2276.

17. V.G.Levich, Physicochemical Hydrodynamics, Prentice-Hall, Englewood Cliffs, 1962, pp. 413-429.

18. T.S.Sørensen, Dynamics and Instability of Fluid Interfaces, Lecture notes on physics, Vol. 276, Springer Verlag, Berlin, 1978.

19. R.Aogaki, K.Kitazawa, K.Fueki and T.Mukaibo, Electrochim. Acta 23 (1978) 867-880.

20. R.G.Barradas and F.Kimmerle, J.Electroanal.Chem. 11 (1966) 163-170.

21. V.Žutić, S.Kovač, J.Tomaić and V.Svetličić, J.Electroanal.Chem. 349 (1993) 173-186.

22. N.Ivošević, J.Tomaić and V.Žutić, Langmuir 10 (1994) 2415-2418.

23. N.Ivošević and V.Žutić, Croat.Chem.Acta 70 (1997) 167-178.

24. S.Kovač, M.Sc.Thesis, Heterocoalescence of Organic Particles at Mercury Electrode in Aqueous Electrolyte Solution and Seawater, University of Zagreb, 1993. (in Croatian)

25. V.Svetličić, N.Ivošević, V,Žutić and D.Fuks, Croat.Chem.Acta 70 (1997) 141-150.

26. R.R.L.Guillard and J.H.Ryter, Com.J.Microbiol. 8 (1962) 229-239.

27. J.Heyrovský and J.Kůta, Principles of Polarography, Czechoslovak Academy of Sciences, Prag, 1965, p 54,81.

28. J.Lyklema and R.Parsons, Electrical Properties of Interfaces. Compilation of Data on the Electrical Double Layer on Mercury Electrodes, Office of Standard Reference Data, National Bureau of Standards, Department of Commerce, Washington, DC, 1983.

29. D.C.Graham, J.Amer.Chem.Soc. 80 (1958) 4201-4210.

30. D.Degobbis, S.Fonda-Umani, P.Franco, A.Malej, R.Precali and N.Smodlaka, Sci.

Tot. Environ. 165 (1995) 43-58.

31. N.Ivošević, M.Sc.Thesis, Organic Particles in the Sea: Interaction with Mercury Electrode, University of Zagreb, 1995. (in Croatian)

32. S.Kovač, V.Svetličić and V.Žutić, Colloids Surf. A: Physicochemical and Engineering Aspects,  149 (1999) 481-489.

33. F.Daly, D.J.Hand, M.C.Jones, A.D.Lunn and K.J.McConway (eds.), Elements of Statistics, Addison-Wesley Publishing Company, 1995, p. 355-358.

34. J.Tomaić, T.Legović and V.Žutić, J.Electroanal.Chem. 322 (1992) 79-92.

35. J.Tomaić, T.Legović and V.Žutić, J.Electroanal.Chem. 259 (1989) 49-57.

36. J.Tomaić, V.Žutić and V.Svetličić, Kem.Ind. 41 (1992) 257-264.

37. S.Kovač and R.Kraus, unpublished results.

SAŽETAK

Suspenzija stanica kao modelni sustav za elektrokemijsku analizu

Solveg Kovač, Romina Kraus, Sunčana Geček i Vera Žutić*

Suspenzija stanica morskog fitoplanktona Dunaliella tertiolecta u otopini 0.1 M NaCl predlaže se kao prikladni monodisperzni modelni sustav (veličina čestica 6-10 (m) za kalibraciju elektrokemijskog odziva u uzorcima prirodnih voda koji sadrže organske površinski-aktivne čestice. Elektrokemijska analiza se izvodi izravnim snimanjem kronoamperometrijskih krivulja redukcije kisika u suspenzijama koristeći brzo-kapajuću živinu elektrodu, koja predstavlja adhezijski senzor. Adhezija pojedinačnih površinski-aktivnih čestica suspendiranih  prisutnih u morskoj vodi, kao  i stanica Dunaliella tertiolecta u modelnim suspenzijama, dovodi do dobro razlučenih signala prianjanja na amperometrijskim krivuljama. Kalibracijska krivulja, koja prikazuje ovisnost učestalosti signala prianjanja o gustoći stanica u rasponu koncentracija 106 do 2.5x107 stanica L-1 može se koristiti za određivanje koncentracije čestica u uzorcima morske vode. Diskutirane su  prednosti elektrokemijskog pristupa pred uobičajenim metodama analize čestica.