MOTION ANALYSIS SYSTEM FOR IDENTIFICATION OF 3D HUMAN LOCOMOTION KINEMATICS DATA AND ACCURACY TESTING

Vladan Papić1, Vlasta Zanchi1

1Faculty of Electrical Engineering, Mechanical Engineering and Naval Architecture,

University of Split, Ruđera Boškovića b.b., 21000 Split, Croatia

vladan.papic@fesb.hr, vlasta.zanchi@fesb.hr

Abstract

In this paper, a procedure for acquisition and processing of the human locomotion system kinematics data is presented. Software for 3D motion analysis system has been developed. Data aquisition was done using two commercial camcorders, framegrabber and PC. Data processing by direct linear transformation provided 3D co-ordinates of the measured movement. System accuracy, regarding error due relatively low frame rate and non-synchronised cameras, was improved using software. Improvement was tested in the laboratory and the results are given. The procedure for synchronisation error elimination will be described in detail.

Keywords

3D motion analysis system, camera synchronisation, direct linear transformation

1	Introduction

Specialised 3D motion systems such as Vicon, Elite, Optotrak (Richards, J., 1998) are becoming standard equipment in gait laboratories and their performances are satisfying the requirements for the clinical gait analysis (Allard, P. et al., 1997). In the perfect world, there will be no serious disadvantages of these movement-tracking systems. Unfortunately, their high-cost is sometimes the most important obstacle when deciding about purchasing new equipment for the gait laboratory. Different family of motion data acquisition systems is based on the video tape and frame-grabber concept (Ariel Life Systems, Peak Performance Technologies Inc.). This approach can also provide real-time data acquisition but its main advantages are in easier manual intervention in resolving possible ambiguities in marker association and easier applying for the outdoors data capturing, such in sports. Generally, because of the less specialised hardware, these systems are cheaper than previously mentioned systems (Vicon, Optotrak). For the normal speed gait, frame rate possibilities of video based systems (25 frames / 50 fields per second) are satisfying the needed number of samples per second. In addition, some new digital cameras (JVC 9800) using Fireware PC connection can provide up to 100 samples per second. Also, a cheaper system could be mounted in larger number of non-specialised laboratories that can be dislocated from the main clinical centres.

There are some critical points considering development of the video tape/framegrabber based 3D motion analysis system. Main problems with this type of system are technical limitations considering system resolution and frame rate, or we can say “the accuracy problems”. Additional problem in obtaining precise 3D marker co-ordinates and reducing differences between two images due to subject motion is synchronisation.

In this paper, a motion analysis system for identification of human locomotion 3D kinematics data using software camera synchronisation will be presented and tested.

2	Measurement and data processing

Whole movement measurement and data processing procedure was carried out in six basic steps. These steps are:

- video tape recording of the moving object in the laboratory

- transferring recorded movement to PC as AVI files for each camera

- extracting calibration co-ordinates

- extracting co-ordinates from the markers positioned on the moving object

- C++ program execution for camera data synchronisation and calculating 3D co-ordinates (DLT)

- graphical presentation of the results.

We used two commercial cameras for the experiment: Sony Digital 8 (DCR-TRV110E) and Panasonic VHS (G101 VHS). Programs were running on Pentium PC with Windows 98 OS. PC was supplied with ATI 3D RAGE framegrabber.

Now, we can describe the procedure step-by-step with special attention on the step 5 in which camera data synchronisation along with data filtering and interpolation is done.

Step 1 - In the laboratory, kinematics data of human locomotion system was measured on a person with attached markers to the right side of the body. Prior to the gait measurement, a calibration cubic frame was recorded by left and right camera. Because of internal orientation calibration, there were no pre-defined positions for the cameras.

Step 2 - Gait was recorded with two cameras simultaneously and avi files for both cameras were obtained. Because of the framegrabber limitations, maximal picture resolution was 640x480 at 25 frames per second.

Step 3 & 4 - Calibration frame co-ordinates were extracted after that, marker co-ordinates for the observed gait cycle were extracted using program developed in C++ Builder. Raw data was saved in the matrix form suitable for further data processing.

Step 5 - camera data synchronisation and 3D co-ordinates calculation. This step demands more detailed explanation. Human gait can be considered as a periodic event, but because of the camera positioning, vertical component of the markers have an offset (in our case, vertical heel co-ordinate in the moment in which heel strikes ground changes proportionally with number of steps).  After taking this fact into consideration, offset could be easily erased using vertical component transformation and data for each marker were filtered using Discrete Fourier Transformation (DFT).

Cut-off frequency was set at 7th harmonic, which is, considering Whittaker-Shannon sampling theorem, high enough for normal walking speed (Winter, D., 1991). Although our frame rate was 25 frames/sec, which means that we had one measured marker co-ordinate every 0.04 s, we have interpolated new data according to the obtained fourier coefficients. For our experiment we have interpolated 100 new points between every measured (and filtered) point but there are no obstacles in increasing that number because computer processing is fast enough. With this new points we obtained one marker co-ordinate every 0.0004 sec, so the highest possible delay between left and right camera data due to lack of ideal hardware synchronisation was 100 times smaller.

Signal maximum for both cameras is the same real moment in time. If we consider one camera (Camera 1) as referent, obtained filtered co-ordinates for the other camera (Camera 2) had to be modified according to delay of frames taken from the Camera 2 and Camera 1.

Lets denote y1,i as the marker co-ordinate y in the moment ti (i-th frame), y1,i+1 co-ordinate y in the moment ti+1 (successive frame) taken from the Camera 1, (t as delay time between two cameras (Fig.1). New, interpolated points between measured marker co-ordinates at ti and ti+1 are added so all values for the Camera 1 can be written as

y1,i+1/n, y1,i+2/n, ..., y1,i+k/n,..., y1,i+1

and for the Camera 2

y2,i+1/n, y2,i+2/n, ..., y2,i+k/n,..., y2,i+1

If the maximum of the signal from the Camera 1 occurs at tj and maximum of the Camera 2 occurs at tj+k/n, then we can calculate the time delay between two cameras

(1)

Transformed co-ordinate y2,i* for the Camera 2 marker data in moment ti is obtained according to equation

y2,i*(ti) = y2,i+k/n()                                                        (2)

Using the same formula we obtained new values for every co-ordinate taken with the Camera 2 in the observed sampling moment. In this way, the cameras are synchronised.

Finally, 3D marker co-ordinates are calculated using the Direct Linear Transformation method (Strintzis, M., & Malassiotis, S., 1999).

|  | Camera (1) |
| --- | --- |
|  | Camera (2) |

Figure 1.

Knee-marker vertical co-ordinate (y) for each camera.

(t denotes time delay between first and second data set.

Step 6 - Processed 3D co-ordinates are visualised using the MATLAB software package (The MathWorks Inc.). Obtained results were used for further gait data analysis (Zanchi, V. et al, 1999). As an example, calculated hip and knee angles presented in 2D are given in Fig. 2.

| a) hip angle | b) knee angle |
| --- | --- |

Figure 2. Calculated hip and knee angles

3	ACCURACY tests

The developed system has been evaluated according to the test protocol used by “The Clinical Gait Analysis Forum of Japan“ (Current Status of 3D Camera systems for Clinical Gait Measurement). Tests were carried-out for various resolutions of the picture (320*240 and 640*480), manual and automatic extraction of the markers and different cut-off frequencies in all three dimensions (Figure 3).

| Figure 3 a. Vertical 90 cm accuracy test | Figure 3 b. Horizontal 90 cm accuracy test |
| --- | --- |
| Figure 3 c. Lateral 90 cm accuracy test | Figure 3 d. Calibration frame |

As it can be seen from the picture, a wooden rod with two 90 cm distant markers was used for the experiment. It was necessary to measure the marker distances as precise as possible prior to the tests. Also, the calibration frame (visible in figure 3 d.) makers were carefully positioned and measured prior to the tests. The 90 cm tests (also called The Separation Test), measured each system's ability to track two points 900 mm apart on a rigid wooden rod while walking through the measurement space.  The test was conducted three times with the rod held in a different orientation on each pass through the space (figures 3 a-c). In addition to the movement in the left to right direction due to the walking of the person holding the person, the rod was moving in up – down direction. This (up-down) movement allowed the software synchronisation of the marker trajectories recorded with both cameras.

The results obtained for the 90 cm accuracy test at 320 x 240 picture resolution are given in table 1. As it can be seen, bias has been almost completely removed and the overall accuracy has been improved. The results for the 640 x 480 resolution provided similar results. It is obvious that the highest error is calculated for the lateral accuracy. It is expected because of the non-quadratic calibration frame. Camera positioning contributed in the high lateral direction error comparing to vertical and progressional direction results.  It can be seen that there is no significant offset error difference for the lateral direction accuracy test. For this case, offset error is rather small in both cases, so it can be said that the synchronisation wasn’t needed because the frame delay between cameras was small. Filtered results are in general more accurate (mean error and standard deviation are smaller).

Comparison between calculated average values of marker distances is given in Figure 4.

| Measurement | Mean value (mm) | Mean abs. error (mm) | Max. error (mm) | Min. error (mm) | Stand. dev. (mm) |
| --- | --- | --- | --- | --- | --- |
| Vertical direction accuracy (picture resolution 320*240) | synchr. | unfiltered | 900.50 | 6.42 | 17.93 | -25.50 | 8.16 |
|  |  | filtered | 900.42 | 5.65 | 13.08 | -15.55 | 6.84 |
|  | unsynchr. | unfiltered | 905.30 | 6.52 | 14.04 | -22.25 | 8.03 |
|  |  | filtered | 905.33 | 5.75 | 10.30 | -16.26 | 6.91 |
| Progressional direction accuracy (picture resolution 320*240) | synchr. | unfiltered | 900.89 | 6.72 | 18.01 | -15.97 | 8.25 |
|  |  | filtered | 901.03 | 6.57 | 17.69 | -19.05 | 7.93 |
|  | unsynchr. | unfiltered | 907.99 | 8.80 | 17.23 | -22.61 | 10.59 |
|  |  | filtered | 908.04 | 8.54 | 17.06 | -27.29 | 10.16 |
| Lateral accuracy (picture resolution 320*240) | synchr. | unfiltered | 899.03 | 9.59 | 30.85 | -33.90 | 12.47 |
|  |  | filtered | 899.03 | 8.57 | 30.84 | -33.89 | 11.16 |
|  | unsynchr. | unfiltered | 899.91 | 10.14 | 32.33 | -34.08 | 13.02 |
|  |  | filtered | 899.87 | 8.78 | 32.43 | -34.02 | 11.43 |

Table 1.

90 cm accuracy test results for the 320*240 picture resolution

Figure 4.

Average values for the filtered data (320 x 240 picture resolution)

| System | No. of Cameras | Mean abs. error (mm) | Max. error (mm) | Min. error (mm) | Stand. dev. (mm) |
| --- | --- | --- | --- | --- | --- |
| Ariel APAS | 2 | 11.61 | 13.47 | -24.07 | 5.36 |
| PEAK5 | 2 | 3.85 | 8.10 | -10.39 | 2.04 |
| VICON 370 | 6 | 0.94 | 4.37 | -8.57 | 0.39 |

Table 2.

Results for Ariel APAS, Peak 5 and Vicon 370 3D movement analysis systems

For the comparison with the existing commercial systems Table 2. presents test results for three commercial systems tested by the Clinical Gait Analysis Forum of Japan. Columns with error values data represents mean error values for all three dimensions.

4	Discussion

The accent was on the accuracy of the observation, so we are focused on improving that factor. At the picture resolution of 640 x 480 pixels and field of view (FOV) approximately 2 m high and 2.7 m wide, error due to quantization is under 0.5 cm, which can be considered as satisfactory. This value is under 1 cm for the 320 x 240 picture resolution. As opposite to this, error due non-synchronised cameras at 25 fps can be 20 times higher (10 cm). This error is calculated for normal walking velocity considering expected highest marker speed. Some manufacturers are trying to solve this problem with synchronisation of the fields but further improvement should be made.

With presented camera synchronisation involving data filtering and interpolation, this error is significantly smaller. As the result of the presented procedure, we have built inexpensive system with the performance within the commercial systems error values. This 3000$ system (two camcorders and PC, software price not included) can be easily mounted and maintained. Although we used only two cameras in this experiment, there are no obstacles for using three or more cameras. More cameras means further accuracy improvement.

References

Allard, P., Capozzo, A., Lundberg, A., ˛& Vaughan, C. (1997). Three-dimensional Analysis of Human Locomotion,  John Wiley & Sons.

Winter, D. (1991). The Biomechanics and Motor Control of Human Gait: Normal Eldery and Pathological, Second  Edition, University of Waterloo Press.

Richards, J., (1998). The Measurement of Human Motion: A Comparison of Commercially Available Systems, Proc. 5th International Symposium on the 3-D Analysis of Human Movement, July 2nd-5th, Chattanooga, Tennessee, USA.

Strintzis, M., Malassiotis, S. (1999). Object-Based Coding of Stereoscopic and 3D Image Sequences, IEEE Signal Processing Magazine, Vol. 16, No.3, pp. 14-28.

Zanchi, V., Cecić, M., & Papić, V. (1999). Visualization of Human Motion Kinematics in State Space, Medical & Biological Engineering & Computing, Vol.37, Supplement 2, Proc. of European Medical & Biological Eng. Conference, EMBEC '99, Vienna, Austria, 4-7 November, Part II, p.p. 846-847.

Zanchi, V., Papić, & V., Zanchi, I. (1999). The Development of Expert System for Identification of Human Walking System, MIC '99, Proceedings of the 18th IASTED International Conference, p.p. 472-474, ISBN: 0-88986-239-7, Innsbruck, Austria, February, 15-18, ACTA Press, Anaheim.

Current Status of 3D Camera systems for Clinical Gait Measurment http://www.aist.go.jp/NIBH/ourpages/comparison99/comp95.html